## Presentation
Abbriviations
MTBF= Mean Time Between Failures
MTTR= Mean Time To Repair
Equations
Beam Availability= (Beam Availability Time)/(Real User Time-InjectionTime)
Ultimate Beam Availability= (Beam Availability Time)/(Real User Time)
Total Down Time= (Failures Time) + (Injection Delay Time)
MTBF= (Beam Availability Time)/(# of Failuers)
MTTR= (FailuersTime)/(# of Failuers)
Machine Mode | (Multiple Items)
Start Time | (Multiple Items)
Row Labels | Sum of Failures Time (min) | Count of Failures
RF | 4208 | 35
PS | 2963 | 4
Cooling | 1727 | 16
Power Grid | 1435 | 7
Electrical | 994 | 2
Control | 869 | 11
Vacuum | 505 | 1
ID | 272 | 6
Microtron | 204 | 7
Diagnostic | 150 | 2
Human Error | 147 | 10
Unknown | 98 | 7
Grand Total | 13572 | 108
SubSystem Failures Time (H) | 226.2 | 13572
Failure time at injection time |  | 1832
No. of failures in user time |  | 53
Beam Availability | 94.43%
Ultimate Beam Availability | 92.9%
MTBF (H) | 68.9664
MTTR (H) | 2.0944
Hours
InjectionTime: | 63.8333
Real UserTime: | 3934.5833
Beam Availability Time: | 3655.2167

## Trip Analysis
Machine Mode | Trip Description | System | Trip Category | Start Time | End Time | Du | Break Down Time | Column1 | Consider | Column2 | Column3
User Time | Shutter Stuck | Front End | Pneumatic | 2020-03-16 16:09 | 2020-03-16 16:37 | 0:28 | 28 | 14400:00 | No
User Time | Beam lost_Unknown | Unknown | Unknown | 2020-06-08 13:03 | 2020-06-08 13:10 | 0:07 | 7 | 14400:00 | Yes
User Time | Shutter Stuck | Front End | Pneumatic | 2020-06-08 19:00 | 2020-06-08 20:00 | 1:00 | 60 | 14400:00 | No
User Time | Server Hanged | Network |  | 2020-06-09 09:30 | 2020-06-09 11:00 | 1:30 | 90 | 14400:00 | No
User Time | Shutter Stuck | Front End | Pneumatic | 2020-06-09 11:24 | 2020-06-09 14:20 | 2:56 | 176 | 14400:00 | No
User Time | Shutter Stuck | Front End | Pneumatic | 2020-06-09 15:41 | 2020-06-09 17:00 | 1:19 | 79 | 14400:00 | No
User Time | SR_BM_PS due to cooling | Cooling |  | 2020-06-09 19:04 | 2020-06-09 19:20 | 0:16 | 16 | 14400:00 | Yes | 13 Min
User Time | Electricity Dip | Power Grid | Power Grid Dip | 2020-06-10 13:51 | 2020-06-10 15:22 | 1:31 | 91 | 14400:00 | Yes
User Time | SR_BM_PS due to cooling | Cooling |  | 2020-06-10 16:57 | 2020-06-10 17:31 | 0:34 | 34 | 14400:00 | Yes
User Time | SR_BM_PS due to cooling | Cooling |  | 2020-06-10 17:57 | 2020-06-10 18:07 | 0:10 | 10 | 14400:00 | Yes | 8 min
Non-Interrupting | MIC_VA | Vacuum | Microtron | 2020-06-11 09:10 | 2020-06-11 09:15 | 0:05 | 5 | 14400:00 | No
User Time | SR_HAB2_C06_CO | Cooling | Flow Switch | 2020-06-14 09:30 | 2020-06-14 10:20 | 0:50 | 50 | 14400:00 | Yes | 3 Min
User Time | Current drop in SR_QD | PS | QD | 2020-06-14 10:50 | 2020-06-14 11:00 | 0:10 | 10 | 14400:00 | Yes
User Time | SR_BM_PS due to cooling | Cooling |  | 2020-06-14 13:10 | 2020-06-14 13:30 | 0:20 | 20 | 14400:00 | Yes
User Time | Beam lost_Human | Human Error | Human Error | 2020-06-14 13:50 | 2020-06-14 13:55 | 0:05 | 5 | 14400:00 | No
User Time | SR_BM_PS due to cooling | Cooling |  | 2020-06-14 14:17 | 2020-06-14 14:45 | 0:28 | 28 | 14400:00 | Yes
User Time | SR_BM_PS due to cooling | Cooling |  | 2020-06-14 16:01 | 2020-06-14 16:11 | 0:10 | 10 | 14400:00 | Yes
Non-Interrupting | MIC_VA | Vacuum | Microtron | 2020-06-16 07:03 | 2020-06-16 07:20 | 0:17 | 17 | 14400:00 | No
User Time | Beam lost_Human | Human Error | Human Error | 2020-06-16 10:24 | 2020-06-16 10:28 | 0:04 | 4 | 14400:00 | Yes | 9 min
User Time | SR_BM_PS due to cooling | Cooling | Chillers | 2020-06-18 10:44 | 2020-06-18 10:51 | 0:07 | 7 | 14400:00 | Yes
Non-Interrupting | MIC_VA | Vacuum | Microtron | 2020-06-18 08:00 | 2020-06-18 08:10 | 0:10 | 10 | 14400:00 | No
Injection | Beam lost_Unknown | Unknown | Unknown | 2020-06-23 09:07 | 2020-06-23 09:21 | 0:14 | 14 | 14400:00 | No
User Time | MIC_VA | Vacuum | Microtron | 2020-06-23 09:44 | 2020-06-23 09:51 | 0:07 | 7 | 14400:00 | No
User Time | SR_BM_PS | PS | SR_BM | 2020-06-23 10:55 | 2020-06-23 13:15 | 2:20 | 140 | 14400:00 | Yes
Injection | Beam lost_Unknown | Unknown | Unknown | 2020-06-25 10:44 | 2020-06-25 10:52 | 0:08 | 8 | 14400:00 | No
Non-Interrupting | MIC_VA | Vacuum | Microtron | 2020-06-25 00:12 | 2020-06-25 00:16 | 0:04 | 4 | 14400:00 | No
Injection | MIC_VA | Vacuum | Microtron | 2020-06-30 09:00 | 2020-06-30 09:30 | 0:30 | 30 | 14400:00 | No
Non-Interrupting | MIC_VA | Vacuum | Microtron | 2020-06-30 09:30 | 2020-06-30 09:45 | 0:15 | 15 | 14400:00 | No
Non-Interrupting | MIC_VA | Vacuum | Microtron | 2020-06-30 10:00 | 2020-06-30 10:10 | 0:10 | 10 | 14400:00 | No
Injection | MIC_VA | Vacuum | Microtron | 2020-07-01 09:44 | 2020-07-01 10:04 | 0:20 | 20 | 14400:00 | No
User Time | SR_BM_PS due to cooling | Cooling | Heat Exchangers | 2020-07-02 13:46 | 2020-07-02 14:00 | 0:14 | 14 | 14400:00 | Yes
User Time | SR_BM couldn't be cycled due to communication error | Control | Communication error | 2020-07-06 09:30 | 2020-07-06 09:50 | 0:20 | 20 | 14400:00 | No
User Time | Arcing in Cavity4 | RF | Cavity4 | 2020-07-06 12:35 | 2020-07-06 12:52 | 0:17 | 17 | 14400:00 | Yes
User Time | Electricity Dip | Power Grid | Power Grid Dip | 2020-07-06 15:44 | 2020-07-06 17:44 | 2:00 | 120 | 14400:00 | Yes
Non-Interrupting | MIC_VA | Vacuum | Microtron | 2020-07-07 08:20 | 2020-07-07 08:40 | 0:20 | 20 | 14400:00 | No
User Time | Beam lost_Unknown | Unknown | Unknown | 2020-07-07 08:46 | 2020-07-07 09:00 | 0:14 | 14 | 14400:00 | Yes
Injection | Booster RF fault | RF | Booster | 2020-07-08 09:00 | 2020-07-08 09:10 | 0:10 | 10 | 14400:00 | No
Injection | Booster RF fault | RF | Booster | 2020-07-09 09:05 | 2020-07-09 09:10 | 0:05 | 5 | 14400:00 | No
Injection | LB2 stuck | Diagnostic | Libera 2 | 2020-07-23 09:25 | 2020-07-23 09:35 | 0:10 | 10 | 14400:00 | No
Injection | Three way valve actuator hanged | Cooling | Primary | 2020-07-23 09:56 | 2020-07-23 10:15 | 0:19 | 19 | 14400:00 | Yes | 6 min
Injection | BO_VA | Vacuum | Ion Pump | 2020-07-27 09:00 | 2020-07-27 09:33 | 0:33 | 33 | 14400:00 | No
User Time | SR_VA_C03 | Vacuum | Gauge Controller | 2020-07-27 13:33 | 2020-07-27 14:04 | 0:31 | 31 | 14400:00 | Yes
User Time | SR_QD2_C06_CO | Cooling | Flow Switch | 2020-07-27 22:11 | 2020-07-28 10:00 | 11:49 | 709 | 14400:00 | Yes | Time entered maniually
Injection | BO_RF_CAV1_REV | RF | Booster | 2020-07-29 09:05 | 2020-07-29 09:06 | 0:01 | 1 | 14400:00 | No
Non-Interrupting | SR_VA_C14_SRM | Vacuum |  | 2020-08-04 08:00 | 2020-08-09 08:00 | 120:00 | 7200 | 14400:00 | No
User Time | SR_VA_C14_SRM | Vacuum |  | 2020-08-09 08:00 | 2020-08-11 17:45 | 57:45 | 3465 | 14400:00 | No
Non-Interrupting | MIC_VA | Vacuum | Microtron | 2020-08-17 03:35 | 2020-08-17 03:45 | 0:10 | 10 | 14400:00 | No
User Time | Electricity Dip | Power Grid | Power Grid Dip | 2020-08-19 04:43 | 2020-08-19 11:30 | 6:47 | 407 | 14400:00 | Yes
Injection | Arcing in Cavity4 | RF | Cavity4 | 2020-08-25 09:29 | 2020-08-25 09:42 | 0:13 | 13 | 14400:00 | Yes
Injection | Beam lost_Human | Human Error | Human Error | 2020-08-25 09:51 | 2020-08-25 09:56 | 0:05 | 5 | 14400:00 | No
User Time | Beam lost_Human | Human Error | Human Error | 2020-08-25 10:56 | 2020-08-25 11:10 | 0:14 | 14 | 14400:00 | Yes
User Time | Arcing in Cavity4 | RF | Cavity4 | 2020-08-26 23:12 | 2020-08-26 23:42 | 0:30 | 30 | 14400:00 | Yes
User Time | SR_QD2_C06_CO | Cooling | Flow Switch | 2020-08-27 13:04 | 2020-08-27 13:53 | 0:49 | 49 | 14400:00 | Yes
User Time | LB2 stuck | Diagnostic | Libera 2 | 2020-09-09 23:16 | 2020-09-10 | 0:44 | 44 | 14400:00 | Yes
User Time | Arcing in Cavity4 | RF | Cavity4 | 2020-09-17 02:09 | 2020-09-17 08:12 | 6:03 | 363 | 14400:00 | Yes
User Time | SR_BM_PS due to cooling | Cooling |  | 2020-09-17 08:35 | 2020-09-17 08:47 | 0:12 | 12 | 14400:00 | Yes
Injection | TL1_Triplet_QD1 | PS | Triplet | 2020-09-21 09:00 | 2020-09-21 10:00 | 1:00 | 60 | 14400:00 | No
User Time | Beam lost_Unknown | Unknown | Unknown | 2020-09-22 08:22 | 2020-09-22 11:41 | 3:19 | 199 | 14400:00 | Yes | Check your self data reate
User Time | Electricity Dip | Power Grid | Power Grid Dip | 2020-09-22 03:55 | 2020-09-22 13:00 | 9:05 | 545 | 14400:00 | Yes
Injection | Shutter Stuck | Front End | Pneumatic | 2020-09-23 09:00 | 2020-09-23 09:02 | 0:02 | 2 | 14400:00 | No
User Time | Shutter Stuck | Front End | Pneumatic | 2020-09-23 16:49 | 2020-09-23 17:12 | 0:23 | 23 | 14400:00 | No
User Time | Arcing in Cavity4 | RF | Cavity4 | 2020-09-23 17:12 | 2020-09-23 17:30 | 0:18 | 18 | 14400:00 | No
User Time | Arcing in Cavity4 | RF | Cavity4 | 2020-10-07 05:59 | 2020-10-07 08:12 | 2:13 | 133 | 14400:00 | Yes
User Time | Arcing in Cavity4 | RF | Cavity4 | 2020-10-08 03:55 | 2020-10-08 08:00 | 4:05 | 245 | 14400:00 | Yes
User Time | Beam lost_Unknown | RF |  | 2020-10-12 11:48 | 2020-10-12 11:55 | 0:07 | 7 | 14400:00 | Yes
Injection | Shutter Stuck | Front End | Pneumatic | 2020-10-20 08:59 | 2020-10-20 09:01 | 0:02 | 2 | 14400:00 | No
Injection | Iearth fault in C11_QF2 | PS | QF | 2020-10-20 09:23 | 2020-10-20 09:45 | 0:22 | 22 | 14400:00 | Yes | 10 Min
Injection | Shutter Stuck | Front End | Pneumatic | 2020-10-21 08:40 | 2020-10-21 08:45 | 0:05 | 5 | 14400:00 | No
Injection | Current drop in C11_QF2 | PS | QF | 2020-10-21 09:05 | 2020-10-21 09:22 | 0:17 | 17 | 14400:00 | Yes
Injection | Shutter Stuck | Front End | Pneumatic | 2020-10-22 09:00 | 2020-10-22 09:09 | 0:09 | 9 | 14400:00 | No
Injection | Current drop in C11_QF2 | PS | QF | 2020-10-22 09:30 | 2020-10-22 09:42 | 0:12 | 12 | 14400:00 | Yes | 9 min
User Time | BM_RC filtration circuit was burned out | PS | SR_BM | 2020-11-17 12:15 | 2020-11-17 16:00 | 3:45 | 225 | 14400:00 | No | No data
User Time | BM_RC filtration circuit was burned out | PS | SR_BM | 2020-11-17 18:53 | 2020-11-17 19:05 | 0:12 | 12 | 14400:00 | No | No data
User Time | BM_RC filtration circuit was burned out | PS | SR_BM | 2020-11-17 20:24 | 2020-11-17 20:54 | 0:30 | 30 | 14400:00 | No | No data
User Time | BM_RC filtration circuit was burned out | PS | SR_BM | 2020-11-17 21:07 | 2020-11-22 14:00 | 112:53 | 2393 | 14400:00 | No | No data
Injection | RF tripped while ramping, due to EPICS gate way didn't give permission to Matlab to write on the LLEs. | Control | EPICS gate way | 2020-11-24 09:18 | 2020-11-24 09:25 | 0:07 | 7 | 14400:00 | No
Injection | RF tripped while ramping, due to EPICS gate way didn't give permission to Matlab to write on the LLEs. | Control | EPICS gate way | 2020-11-24 09:36 | 2020-11-24 10:49 | 1:13 | 73 | 14400:00 | No
Injection | Beam lost; due to BPM interlock. | Diagnostic |  | 2020-11-24 11:07 | 2020-11-24 11:20 | 0:13 | 13 | 14400:00 | Yes
Injection | Beam lost; due to BPM interlock. | Diagnostic |  | 2020-11-24 11:40 | 2020-11-24 11:51 | 0:11 | 11 | 14400:00 | Yes
Injection | Beam lost; due to BPM interlock. | Diagnostic |  | 2020-11-24 12:12 | 2020-11-24 13:26 | 1:14 | 74 | 14400:00 | Yes
Injection | Beam lost; due to BPM interlock. | Diagnostic |  | 2020-11-24 13:48 | 2020-11-24 13:52 | 0:04 | 4 | 14400:00 | Yes
Injection | Beam lost; due to BPM interlock. | Diagnostic |  | 2020-11-24 14:21 | 2020-11-24 14:33 | 0:12 | 12 | 14400:00 | Yes
Injection | Iearth fault in C11_QF2 | PS | QF | 2020-12-01 09:32 | 2020-12-01 09:54 | 0:22 | 22 | 14400:00 | Yes
Injection | Iearth fault in C11_QF2 | PS | QF | 2020-12-01 10:13 | 2020-12-01 10:26 | 0:13 | 13 | 14400:00 | Yes
Injection | Iearth fault in C11_QF2 | PS | QF | 2020-12-01 10:50 | 2020-12-01 11:06 | 0:16 | 16 | 14400:00 | Yes
Injection | Iearth fault in C11_QF2 | PS | QF | 2020-12-01 11:29 | 2020-12-01 12:25 | 0:56 | 56 | 14400:00 | Yes
Injection | Iearth fault in C11_QF2 | PS | QF | 2020-12-01 12:48 | 2020-12-01 14:06 | 1:18 | 78 | 14400:00 | Yes
User Time | Current drop in C14_QF1 | PS | QF | 2020-12-02 10:21 | 2020-12-02 14:14 | 3:53 | 233 | 14400:00 | Yes
User Time | SR_BM_PS | PS | SR_BM | 2020-12-02 12:39 | 2020-12-02 17:34 | 4:55 | 295 | 14400:00 | No
User Time | Beam killed by MPS (SR-MAIN OPERATION KEY OFF by mistake) | Human Error |  | 2020-12-05 14:04 | 2020-12-05 14:08 | 0:04 | 4 | 14400:00 | No
Non-Interrupting | Filament couldn't run, after changing the Aux. Gun | Human Error |  | 2020-12-08 08:30 | 2020-12-08 14:00 | 5:30 | 330 | 14400:00 | No
User Time | Electricity Dip | Power Grid |  | 2020-12-08 21:56 | 2020-12-09 10:00 | 12:04 | 64 | 14400:00 | Yes
User Time | SR_VA_C01 | Vacuum |  | 2020-12-09 18:10 | 2020-12-09 18:24 | 0:14 | 14 | 14400:00 | Yes
User Time | SR_BM_PS | PS | SR_BM | 2020-12-15 08:45 | 2020-12-15 15:22 | 6:37 | 397 | 14400:00 | Yes
User Time | RF Interlock | RF |  | 2020-12-15 16:14 | 2020-12-15 16:26 | 0:12 | 12 | 14400:00 | Yes
User Time | LB2 stuck | Diagnostic | Libera 2 | 2020-12-16 17:02 | 2020-12-16 17:11 | 0:09 | 9 | 14400:00 | Yes
Free | Current drop in multiple QFs & QDs | PS | QF & QD | 2020-12-18 15:20 | 2020-12-19 12:02 | 20:42 | 1242 | 14400:00 | Yes
User Time | Klysron-Phase Correction | RF | Klystron Phase Correction | 2020-12-20 12:50 | 2020-12-20 12:58 | 0:08 | 8 | 14400:00 | No
Non-Interrupting | Current drop in multiple QFs & QDs | PS | QF & QD | 2020-12-23 13:16 | 2020-12-23 13:20 | 0:04 | 4 | 14400:00 | No
User Time | SR_VA_C08 | Vacuum |  | 2020-12-24 22:14 | 2020-12-24 22:20 | 0:06 | 6 | 14400:00 | Yes
Non-Interrupting | BO_RF Interlock | Control | Power Supplies | 2021-02-08 15:00 | 2021-02-08 15:10 | 0:10 | 10 | 14400:00 | No
User Time | SR_BM_PS | PS | SR_BM | 2021-02-09 12:56 | 2021-02-09 13:20 | 0:24 | 24 | 14400:00 | Yes
Injection | Beam lost_Unknown | Unknown |  | 2021-02-11 09:45 | 2021-02-11 09:55 | 0:10 | 10 | 14400:00 | Yes
User Time | Electricity Dip | Power Grid |  | 2021-02-16 12:07 | 2021-02-16 13:14 | 1:07 | 67 | 14400:00 | Yes
User Time | Iearth fault in C04_QD2 | PS |  | 2021-02-24 10:55 | 2021-02-24 11:10 | 0:15 | 15 | 14400:00 | Yes
User Time | RF Interlock | RF |  | 2021-02-24 19:22 | 2021-02-24 20:00 | 0:38 | 38 | 14400:00 | Yes | Time entered maniually
Injection | Wave guide pressure fault | RF |  | 2021-02-27 09:00 | 2021-02-27 10:55 | 1:55 | 115 | 14400:00 | No
Injection | Beam lost_due to FILK system | Diagnostic |  | 2021-03-02 09:15 | 2021-03-02 09:25 | 0:10 | 10 | 14400:00 | No | 1 min
Injection | Beam lost_due to FILK system | Diagnostic |  | 2021-03-03 08:20 | 2021-03-03 12:20 | 4:00 | 240 | 14400:00 | Yes
Machine Time | SR_BM_C07_CO | Cooling |  | 2021-03-14 10:00 | 2021-03-14 10:10 | 0:10 | 10 | 14400:00 | Yes
User Time | XAF Beam line closed the gate valve 1 by mistake | Human Error |  | 2021-03-16 09:30 | 2021-03-16 09:42 | 0:12 | 12 | 14400:00 | No
User Time | Reverse power SSA1-DIS09 | RF | SSA1 | 2021-03-20 09:11 | 2021-03-20 11:03 | 1:52 | 112 | 14400:00 | Yes | 4 min
User Time | Reverse power SSA1-DIS09 | RF |  | 2021-03-21 09:35 | 2021-03-21 10:11 | 0:36 | 36 | 14400:00 | Yes | 4 min
User Time | Reverse power SSA1-DIS09 | RF |  | 2021-03-22 09:44 | 2021-03-22 10:15 | 0:31 | 31 | 14400:00 | Yes
Non-Interrupting | Emission Fault | Vacuum |  | 2021-03-22 18:40 | 2021-03-22 19:33 | 0:53 | 53 | 14400:00 | No
Injection | RF Went inside SR as restricted acees for checking subsystem. | RF |  | 2021-03-23 09:02 | 2021-03-23 09:16 | 0:14 | 14 | 14400:00 | No
Injection | RF replaced some module in SSA | RF |  | 2021-03-29 09:02 | 2021-03-29 09:29 | 0:27 | 27 | 14400:00 | No
Injection | Tuners 2 and 3 | RF |  | 2021-04-01 08:21 | 2021-04-01 08:46 | 0:25 | 25 | 14400:00 | No
User Time | Emission Fault | Microtron |  | 2021-04-13 09:26 | 2021-04-13 15:00 | 5:34 | 334 | 14400:00 | No
User Time | Vacuum Leakage after changing the aux. gun | Vacuum | Microtron | 2021-04-13 15:00 | 2021-04-14 17:24 | 26:24 | 1584 | 600 | No
User Time | SR_BM_C07_CO | Cooling |  | 2021-04-14 18:17 | 2021-04-15 11:00 | 16:43 | 1003 | 600 | Yes
Free | Beam lost due to BPM interlock | Diagnostic |  | 2021-04-16 14:26 | 2021-04-17 09:00 | 18:34 | 1114 | 14400:00 | Yes
User Time | SR_BM_PS | PS | SR_BM | 2021-04-17 11:50 | 2021-04-17 12:10 | 0:20 | 20 | 14400:00 | Yes
User Time | Beam lost due to RF trip | RF |  | 2021-04-19 09:33 | 2021-04-19 13:05 | 3:32 | 212 | 14400:00 | No
User Time | Beam lost due to RF trip | RF |  | 2021-04-19 16:00 | 2021-04-19 16:16 | 0:16 | 16 | 14400:00 | Yes
User Time | Beam lost due to loose connection of the cable in the PLC communication module | Control | Communication error | 2021-04-20 14:13 | 2021-04-20 15:16 | 1:03 | 63 | 14400:00 | No
User Time | Beam lost due to human error | Human Error |  | 2021-04-20 15:11 | 2021-04-20 15:15 | 0:04 | 4 | 14400:00 | No
User Time | Bypass the fast interlock system | Diagnostic |  | 2021-04-25 09:17 | 2021-04-25 09:41 | 0:24 | 24 | 14400:00 | No
User Time | SR_BM_PS | PS | SR_BM | 2021-04-25 13:16 | 2021-04-25 13:48 | 0:32 | 32 | 14400:00 | Yes
User Time | Beam lost_Unknown | RF |  | 2021-04-26 01:55 | 2021-04-26 09:00 | 7:05 | 425 | 360 | Yes
User Time | SR_BM_CROWBAR interlock | PS | SR_BM | 2021-05-20 14:40 | 2021-05-20 16:00 | 1:20 | 80 | 14400:00 | Yes
Free | Electricity Dip | Power Grid |  | 2021-05-21 11:45 | 2021-05-21 11:50 | 0:05 | 5 | 14400:00 | Yes
User Time | RF oscillation | RF |  | 2021-05-27 09:26 | 2021-05-27 09:46 | 0:20 | 20 | 14400:00 | No
User Time | RF oscillation | RF |  | 2021-05-30 09:22 | 2021-05-30 09:31 | 0:09 | 9 | 14400:00 | No
User Time | Electricity Dip | Power Grid |  | 2021-05-31 16:51 | 2021-05-31 17:52 | 1:01 | 61 | 14400:00 | Yes
User Time | Chiller was in Local instead of Remote mode after electricity dip | Human Error |  | 2021-05-31 20:28 | 2021-05-31 20:40 | 0:12 | 12 | 14400:00 | Yes
User Time | RF oscillation | RF |  | 2021-06-03 09:25 | 2021-06-03 09:32 | 0:07 | 7 | 14400:00 | Yes
User Time | RF oscillation | RF |  | 2021-06-07 09:29 | 2021-06-07 09:37 | 0:08 | 8 | 14400:00 | No
Injection | LLE1 Reflected Power | RF |  | 2021-06-22 09:11 | 2021-06-22 09:16 | 0:05 | 5 | 14400:00 | No
User Time | SSA1 Water Flow | RF |  | 2021-06-22 09:45 | 2021-06-22 10:00 | 0:15 | 15 | 14400:00 | Yes | 15 Min
Injection | Libera7 disconnection | Diagnostic |  | 2021-06-23 09:35 | 2021-06-23 09:47 | 0:12 | 12 | 14400:00 | Yes
User Time | Klysron-Phase Correction | RF | Klystron Phase Correction | 2021-06-23 10:28 | 2021-06-23 10:41 | 0:13 | 13 | 14400:00 | Yes
User Time | Electricity Dip | Power Grid | Power Grid Dip | 2021-06-29 14:11 | 2021-06-29 14:27 | 0:16 | 16 | 14400:00 | Yes
User Time | Glitch in the Q3 Fast interlock | Diagnostic | Fast Interlock | 2021-06-30 14:41 | 2021-06-30 15:06 | 0:25 | 25 | 14400:00 | Yes
User Time | Beam lost_Unknown | Unknown |  | 2021-07-01 09:42 | 2021-07-01 09:54 | 0:12 | 12 | 14400:00 | No
User Time | Flow Switch of RF Cooling Racks | RF |  | 2021-07-10 15:05 | 2021-07-10 16:38 | 1:33 | 93 | 14400:00 | Yes
User Time | SR_HAB4_C05_CO | Cooling | Heat Absorber | 2021-07-10 17:05 | 2021-07-10 17:30 | 0:25 | 25 | 14400:00 | Yes
Injection | Wiggler Downstream Kill Switch | ID | Wiggler | 2021-08-17 08:05 | 2021-08-17 09:19 | 1:14 | 74 | 14400:00 | No
Non-Interrupting | Booster Dipole PS Interlock | PS | BO_BM | 2021-08-18 09:00 | 2021-08-18 10:10 | 1:10 | 70 | 14400:00 | No
Non-Interrupting | HVPS Interlock | PS | Modulator | 2021-08-19 09:00 | 2021-08-19 14:25 | 5:25 | 325 | 14400:00 | No
User Time | SR_VA_IMG2_C03 | Vacuum |  | 2021-08-21 12:08 | 2021-08-21 12:17 | 0:09 | 9 | 14400:00 | Yes
Machine Time | Wiggler Calibration | ID | Wiggler | 2021-08-22 09:00 | 2021-08-22 12:12 | 3:12 | 192 | 14400:00 | Yes
User Time | SR_BM_C09_CO | Cooling |  | 2021-08-24 01:07 | 2021-08-24 10:50 | 9:43 | 583 | 420 | Yes
User Time | SR_BM_C12_CO | Cooling |  | 2021-08-30 20:30 | 2021-08-31 09:45 | 13:15 | 795 | 705 | Yes
User Time | RF Interlock | RF |  | 2021-08-31 10:16 | 2021-08-31 10:23 | 0:07 | 7 | 14400:00 | Yes
User Time | Beam lost_Human | Human Error |  | 2021-09-02 19:14 | 2021-09-02 22:40 | 3:26 | 206 | 14400:00 | Yes
Non-Interrupting | Wiggler Downstream Kill Switch | ID | Wiggler | 2021-09-02 21:14 | 2021-09-02 22:40 | 1:26 | 86 | 14400:00 | No
User Time | Electricity Dip | Power Grid | Power Grid Dip | 2021-09-03 11:50 | 2021-09-03 14:38 | 2:48 | 168 | 14400:00 | Yes
User Time | Beam lost while ramping due to RF voltage stopped increasing with SR-BM current (ramping script) | Human Error |  | 2021-09-23 10:42 | 2021-09-23 10:46 | 0:04 | 4 | 14400:00 | Yes | 3 Min
User Time | Beam lost due to arcing in cavity 4 | RF | Cavity4 | 2021-09-23 19:01 | 2021-09-23 19:18 | 0:17 | 17 | 14400:00 | Yes
Machine Time | Beam lost due to vacuum in the cavity 4 | RF | Cavity4 | 2021-09-25 18:55 | 2021-09-25 18:55 | 0:00 | 0 | 14400:00 | Yes
Machine Time | Interlock in C04_QD2 | PS | QD | 2021-09-25 20:21 | 2021-09-25 20:21 | 0:00 | 0 | 14400:00 | Yes
Machine Time | Beam lost due to arcing in cavity 4 | RF | Cavity4 | 2021-09-25 21:08 | 2021-09-25 21:08 | 0:00 | 0 | 14400:00 | Yes
Machine Time | Beam lost due to arcing in cavity 4 | RF | Cavity4 | 2021-09-25 22:12 | 2021-09-25 22:12 | 0:00 | 0 | 14400:00 | Yes
User Time | MPS interlock at IR shutter | Front End |  | 2021-09-28 08:25 | 2021-09-28 08:40 | 0:15 | 15 | 14400:00 | Yes
User Time | Beam lost due to arcing in cavity 4 | RF | Cavity4 | 2021-09-29 04:30 | 2021-09-29 08:16 | 3:46 | 226 | 226 | Yes
User Time | Beam lost due to arcing in cavity 4 | RF | Cavity4 | 2021-10-06 07:11 | 2021-10-06 07:55 | 0:44 | 44 | 44 | Yes
User Time | SR_BM_PS_PCO | PS | PCO | 2021-10-07 13:34 | 2021-10-07 13:59 | 0:25 | 25 | 14400:00 | Yes
User Time | XAF Beam line closed the gate valve 1 by mistake | Human Error |  | 2021-10-07 17:37 | 2021-10-07 17:50 | 0:13 | 13 | 14400:00 | No
Machine Time | Beam Lost due to cable loose at RF system and it gives fake vacuum signal | RF |  | 2021-10-20 08:56 | 2021-10-20 09:05 | 0:09 | 9 | 14400:00 | Yes
Machine Time | Electricity Dip | Power Grid | Power Grid Dip | 2021-10-20 09:37 | 2021-10-20 11:28 | 1:51 | 111 | 14400:00 | Yes
Machine Time | SR_BM_PS_PCO | PS | PCO | 2021-10-20 12:38 | 2021-10-20 18:40 | 6:02 | 362 | 14400:00 | Yes
User Time | Beam lost due to arcing in cavity 4 | RF | Cavity4 | 2021-10-25 08:44 | 2021-10-25 08:55 | 0:11 | 11 | 14400:00 | Yes | 3 Min
User Time | SR_BM_PS_PCO | PS | PCO | 2021-10-25 14:33 | 2021-10-25 16:55 | 2:22 | 142 | 14400:00 | Yes
User Time | SR_BM_PS_PCO | PS | PCO | 2021-10-26 10:45 | 2021-10-26 19:42 | 8:57 | 537 | 14400:00 | Yes
User Time | Beam lost due to Chiller 2 was off since starting of the morning injection | Human Error |  | 2021-10-28 10:21 | 2021-10-28 10:42 | 0:21 | 21 | 14400:00 | Yes
Injection | Beam lost due to arcing in cavity 4 | RF | Cavity4 | 2021-11-03 09:28 | 2021-11-03 09:40 | 0:12 | 12 | 14400:00 | Yes
User Time | Beam lost due to loosen connection on the magnet side | Magnet |  | 2021-11-07 05:59 | 2021-11-07 08:43 | 2:44 | 164 | 163 | Yes
Injection | Beam lost due to cavity 2 limit switch was active | RF | Cavity4 | 2021-11-07 08:43 | 2021-11-07 08:59 | 0:16 | 16 | 14400:00 | No
Injection | Beam lost due to RF test | Human Error |  | 2021-11-09 08:38 | 2021-11-09 08:41 | 0:03 | 3 | 14400:00 | No
Injection | Beam lost due to Libera4-BPM disconnection | Human Error |  | 2021-11-10 08:49 | 2021-11-10 08:59 | 0:10 | 10 | 14400:00 | Yes | 3 Min
User Time | Electricity Dip | Power Grid | Power Grid Dip | 2021-11-11 22:18 | 2021-11-12 08:00 | 9:42 | 582 | 582 | Yes
User Time | Phase failure in quadrant#1 sub-main distribution board | Power Grid | Power Grid Dip | 2021-11-22 03:15 | 2021-11-22 12:20 | 9:05 | 545 | 285 | Yes
User Time | Beam lost_Unknown | Unknown | Unknown | 2021-11-22 17:07 | 2021-11-22 17:30 | 0:23 | 23 | 14400:00 | Yes
User Time | AGPS Gun sparking | PS | Aux. Gun | 2021-11-27 08:15 | 2021-11-27 09:20 | 1:05 | 65 | 14400:00 | No
User Time | TL2 magnet & Microtron dipole PS communication Error | PS | Aux. Gun | 2021-11-28 09:00 | 2021-11-28 11:12 | 2:12 | 132 | 14400:00 | No
User Time | Beam lost_Unknown | Unknown | Unknown | 2021-11-28 11:12 | 2021-11-28 11:23 | 0:11 | 11 | 14400:00 | Yes | 2 min
User Time | Beam lost_Unknown | Unknown | Unknown | 2021-11-30 03:37 | 2021-11-30 08:26 | 4:49 | 289 | 263 | Yes
User Time | Klysron-Phase Correction | RF | Klystron Phase Correction | 2021-11-30 08:52 | 2021-11-30 09:26 | 0:34 | 34 | 14400:00 | Yes | Time entered maniually
User Time | Klysron-Phase Correction | RF | Klystron Phase Correction | 2021-12-01 08:59 | 2021-12-01 09:13 | 0:14 | 14 | 14400:00 | Yes
User Time | AGPS Gun sparking | PS | Aux. Gun | 2021-12-02 08:20 | 2021-12-02 08:35 | 0:15 | 15 | 14400:00 | No
User Time | RF group did small test in service area | RF |  | 2021-12-02 08:35 | 2021-12-02 08:56 | 0:21 | 21 | 14400:00 | No
User Time | Beam lost_Unknown | Unknown | Unknown | 2021-12-07 15:38 | 2021-12-07 16:23 | 0:45 | 45 | 14400:00 | Yes
User Time | Beam lost due to RF interlock | RF |  | 2021-12-13 12:23 | 2021-12-13 12:41 | 0:18 | 18 | 14400:00 | Yes
User Time | Beam lost_Unknown | Unknown | Unknown | 2021-12-13 19:08 | 2021-12-13 19:23 | 0:15 | 15 | 14400:00 | Yes
User Time | Beam lost_Unknown | Unknown | Unknown | 2021-12-14 08:19 | 2021-12-14 09:13 | 0:54 | 54 | 14400:00 | Yes
User Time | AGPS Gun sparking | PS | Aux. Gun | 2021-12-15 08:20 | 2021-12-15 08:24 | 0:04 | 4 | 14400:00 | No
User Time | Electricity Dip | Power Grid | Power Grid Dip | 2021-12-15 08:32 | 2021-12-15 08:43 | 0:11 | 11 | 14400:00 | Yes
User Time | Beam lost_Human | Human Error |  | 2021-12-15 10:33 | 2021-12-15 10:48 | 0:15 | 15 | 14400:00 | No
User Time | Klysron-Phase Correction | RF | Klystron Phase Correction | 2021-12-15 11:07 | 2021-12-15 11:16 | 0:09 | 9 | 14400:00 | Yes | 3 Min
User Time | AGPS Gun sparking | PS | Aux. Gun | 2021-12-15 11:26 | 2021-12-15 11:43 | 0:17 | 17 | 14400:00 | No
User Time | Klysron-Phase Correction | RF | Klystron Phase Correction | 2021-12-16 08:18 | 2021-12-16 08:55 | 0:37 | 37 | 14400:00 | Yes
User Time | Beam lost due to RF interlock | RF |  | 2021-12-22 08:54 | 2021-12-22 09:06 | 0:12 | 12 | 14400:00 | Yes
Machine Time | Beam lost due to RF interlock | RF |  | 2022-02-16 06:50 | 2022-02-16 08:15 | 1:25 | 85 | 14400:00 | Yes
Machine Time | Beam lost due to RF interlock | RF |  | 2022-02-20 18:35 | 2022-02-20 18:40 | 0:05 | 5 | 14400:00 | Yes
Machine Time | Beam lost due to RF interlock | RF |  | 2022-02-20 19:10 | 2022-02-20 19:15 | 0:05 | 5 | 14400:00 | Yes
Machine Time | Beam lost due to RF interlock | RF |  | 2022-02-21 05:51 | 2022-02-21 08:13 | 2:22 | 142 | 14400:00 | Yes
User Time | Beam lost due to slow ADC channel 0 | RF |  | 2022-02-21 08:40 | 2022-02-21 08:47 | 0:07 | 7 | 14400:00 | Yes
User Time | Beam lost due to Vacuum in the cavity 3 | RF | Cavity3 | 2022-02-21 12:17 | 2022-02-21 12:42 | 0:25 | 25 | 14400:00 | Yes
User Time | Beam lost due to Vacuum in the cavity 3 | RF | Cavity3 | 2022-02-23 09:05 | 2022-02-23 09:24 | 0:19 | 19 | 14400:00 | Yes
User Time | Beam lost due to Vacuum in the cavity 4 | RF | Cavity4 | 2022-02-23 09:48 | 2022-02-23 09:52 | 0:04 | 4 | 14400:00 | Yes
Machine Time | Beam lost due to Vacuum in the cavity 4 | RF | Cavity4 | 2022-02-26 09:30 | 2022-02-26 09:40 | 0:10 | 10 | 14400:00 | Yes
Machine Time | Beam lost due to Vacuum in the cavity 4 | RF | Cavity4 | 2022-02-26 23:59 | 2022-02-26 23:59 | 0:00 | 0 | 14400:00 | Yes
Machine Time | Beam lost due to RF interlock | RF |  | 2022-03-03 11:53 | 2022-03-03 12:02 | 0:09 | 9 | 14400:00 | Yes
Machine Time | Beam lost due to RF interlock | RF |  | 2022-03-03 12:16 | 2022-03-03 12:26 | 0:10 | 10 |  | Yes
Machine Time | Beam lost due to Vacuum in the cavity 3 | RF | Cavity3 | 2022-03-03 12:46 | 2022-03-03 12:51 | 0:05 | 5 | 14400:00 | Yes
Machine Time | Beam lost due to slow ADC channel 2 | RF |  | 2022-03-05 12:37 | 2022-03-05 12:43 | 0:06 | 6 |  | Yes
Machine Time | Beam lost due to Vacuum in cell #8 | Vacuum |  | 2022-03-07 06:59 | 2022-03-07 08:24 | 1:25 | 85 | 14400:00 | Yes
User Time | Beam lost due to vacuum interlock at cell #8 | Vacuum |  | 2022-03-07 08:00 | 2022-03-07 08:24 | 0:24 | 24 | 14400:00 | No | Unknown recorded at 7:00
User Time | Beam lost due to Vacuum in the cavity 3 | RF | Cavity3 | 2022-03-07 08:43 | 2022-03-07 08:47 | 0:04 | 4 | 14400:00 | No | 6 min
User Time | Electricity Dip | Power Grid | Power Grid Dip | 2022-03-08 17:41 | 2022-03-09 10:50 | 17:09 | 1029 | 14400:00 | Yes
User Time | Beam lost due some changes in RF parameters after the power cut (Cavity #1 forward power high limit was 60 instead of 85) | Power Grid |  | 2022-03-09 10:50 | 2022-03-09 12:44 | 1:54 | 114 | 14400:00 | No
User Time | Beam lost due to Forward power in Cavity 1&3 | RF |  | 2022-03-09 16:09 | 2022-03-09 17:37 | 1:28 | 88 | 14400:00 | Yes
User Time | Beam lost due to Vacuum in the cavity 3 | RF | Cavity3 | 2022-03-13 09:23 | 2022-03-13 09:29 | 0:06 | 6 | 14400:00 | Yes |  | Added by AB MZ
User Time | Beam lost due to Vacuum in the cavity 3 | RF | Cavity3 | 2022-03-13 08:46 | 2022-03-13 09:29 | 0:43 | 43 | 14400:00 | No | 5 min
User Time | Beam lost due to slow ADC channel 0 | RF |  | 2022-03-13 10:18 | 2022-03-13 10:36 | 0:18 | 18 | 14400:00 | Yes
User Time | AGPS Filament OverVoltage | PS | Aux. Gun | 2022-03-13 10:36 | 2022-03-13 10:43 | 0:07 | 7 | 14400:00 | No
User Time | Beam lost due to Vacuum in the cavity 3 | RF | Cavity3 | 2022-03-13 13:07 | 2022-03-13 13:25 | 0:18 | 18 | 14400:00 | Yes
User Time | Beam lost due to RF interlock happened after changing some RF parameters mistakenly | Human Error |  | 2022-03-13 13:35 | 2022-03-13 13:40 | 0:05 | 5 | 14400:00 | No | 5 min
User Time | Beam lost due to Vacuum in cell #8 | Vacuum |  | 2022-03-13 23:00 | 2022-03-13 23:25 | 0:25 | 25 | 14400:00 | Yes
User Time | Beam lost due to Vacuum in the cavity 3 | RF | Cavity3 | 2022-03-14 08:28 | 2022-03-14 08:35 | 0:07 | 7 | 14400:00 | No | 3 min
User Time | Beam lost due to Vacuum in the cavity 4 | RF | Cavity4 | 2022-03-14 12:13 | 2022-03-14 12:20 | 0:07 | 7 | 14400:00 | Yes
User Time | Beam lost due to Vacuum in the cavity 4 | RF | Cavity4 | 2022-03-14 13:53 | 2022-03-14 19:06 | 5:13 | 313 | 14400:00 | Yes
User Time | Beam lost due to insufficient air; due to both of the air compressors turned off due to human mistake | Human Error |  | 2022-03-15 08:26 | 2022-03-15 10:05 | 1:39 | 99 | 14400:00 | No | 7 min
User Time | Beam lost due to electricity dip | Power Grid |  | 2022-03-18 06:48 | 2022-03-18 08:00 | 1:12 | 72 | 14400:00 | Yes
User Time | flow switch interlock on the MS beamline front end | Cooling | Flow Switch | 2022-03-21 08:10 | 2022-03-21 08:58 | 0:48 | 48 | 14400:00 | Yes
User Time | Beam lost due to check valve of RF cooling water pump was clogged | Cooling | RF Secondary Pumps | 2022-03-23 09:04 | 2022-03-23 09:48 | 0:44 | 44 | 14400:00 | Yes
User Time | Beam lost_Unknown | Unknown | Unknown | 2022-03-24 11:44 | 2022-03-24 11:50 | 0:06 | 6 | 14400:00 | Yes
Machine Time | Beam lost due to Vacuum in RF cavitities | RF |  | 2022-03-26 05:49 | 2022-03-26 11:55 | 6:06 | 366 | 14400:00 | Yes
User Time | AGPS Filament OverVoltage | PS | Aux. Gun | 2022-03-28 08:30 | 2022-03-28 08:38 | 0:08 | 8 | 14400:00 | No
User Time | Beam lost due to BPM2 at cell #15 (BPM number 58) exceeded the orbit interlock limit | Diagnostic | Beam Position Monitor | 2022-03-30 09:46 | 2022-03-30 09:55 | 0:09 | 9 | 14400:00 | Yes
User Time | AGPS Filament OverVoltage | PS | Aux. Gun | 2022-03-30 10:03 | 2022-03-30 10:11 | 0:08 | 8 | 14400:00 | No
User Time | Beam lost due to cavities interaction with beam instability | RF |  | 2022-04-01 05:10 | 2022-04-01 08:00 | 2:50 | 170 | 14400:00 | Yes
Non-Interrupting | Wave guide pressure fault | RF | Booster | 2022-04-03 08:00 | 2022-04-03 09:00 | 1:00 | 60 | 14400:00 | No
Beamline Preparation | SR_VA_C01 | Vacuum |  | 2022-06-07 14:58 | 2022-06-07 17:11 | 2:13 | 133 | 14400:00 | Yes | Possible noise
Beamline Preparation | Current limit interlock in QD | PS | QD | 2022-06-09 15:38 | 2022-06-09 16:40 | 1:02 | 62 | 14400:00 | Yes | Possible noise
User Time | Beam lost while bumping, due to BPM interlock was enabled | Human Error |  | 2022-06-15 19:18 | 2022-06-15 19:21 | 0:03 | 3 | 14400:00 | Yes
User Time | Correction was not running properly | Unknown |  | 2022-06-15 21:25 | 2022-06-15 21:30 | 0:05 | 5 | 14400:00 | Yes
User Time | Beam Lost after enabling the BPM directly. We figured out that the min X limit was 1mm instead of -1mm. | Human Error |  | 2022-06-15 22:36 | 2022-06-15 22:40 | 0:04 | 4 | 14400:00 | Yes
User Time | Beam lost while increasing RF voltage to 400 KV, reflected power. | RF | Klystron Phase Correction | 2022-06-15 23:49 | 2022-06-16 00:26 | 0:37 | 37 | 14400:00 | Yes
Beamline Preparation | MS beamline scientist asked to enter restricted access to SR with vacuum group to install rotating fillter | Front End |  | 2022-06-20 11:41 | 2022-06-20 17:46 | 6:05 | 365 | 14400:00 | No
Beamline Preparation | SR_VA_C01 | Vacuum |  | 2022-06-18 02:49 | 2022-06-18 09:00 | 6:11 | 371 | 14400:00 | Yes | Time entered maniually
Beamline Preparation | AGPS Filament OverVoltage | PS | Aux. Gun | 2022-06-18 09:00 | 2022-06-18 12:08 | 3:08 | 188 | 14400:00 | No
Beamline Preparation | Beam got lost due to RF interlock | RF |  | 2022-06-18 17:21 | 2022-06-18 22:41 | 5:20 | 320 | 14400:00 | Yes
Beamline Preparation | AGPS Filament OverVoltage | PS | Aux. Gun | 2022-06-18 22:41 | 2022-06-18 22:58 | 0:17 | 17 | 14400:00 | No
User Time | Electricity Dip | Power Grid |  | 2022-06-22 15:53 | 2022-06-22 16:23 | 0:30 | 30 | 14400:00 | Yes
User Time | SR_VA_C08 | Vacuum |  | 2022-06-23 04:48 | 2022-06-23 08:46 | 3:58 | 238 | 14400:00 | Yes
Beamline Preparation | MS beamline scientist asked to enter restricted access to SR with vacuum group to install rotating fillter | Front End |  | 2022-06-26 10:37 | 2022-06-26 16:37 | 6:00 | 360 | 14400:00 | No
User Time | Beam lost due to Cavity 3 Interlock | RF | Klystron Phase Correction | 2022-06-26 16:44 | 2022-06-26 16:52 | 0:08 | 8 | 14400:00 | No
Beamline Preparation | MS beamline scientist asked to enter restricted access to SR with vacuum group to install rotating fillter | Front End |  | 2022-06-27 09:27 | 2022-06-27 13:54 | 4:27 | 267 | 14400:00 | No
Beamline Preparation | MS beamline scientist asked to enter restricted access to SR with vacuum group to install rotating fillter | Front End |  | 2022-06-28 09:38 | 2022-06-28 10:42 | 1:04 | 64 | 14400:00 | No
Beamline Preparation | MS beamline scientist asked to enter restricted access to SR with vacuum group to install rotating fillter | Front End |  | 2022-06-28 15:30 | 2022-06-28 16:11 | 0:41 | 41 | 14400:00 | No
User Time | Beam lost due to BPM interlock. | Diagnostic |  | 2022-06-30 18:05 | 2022-06-30 18:21 | 0:16 | 16 | 14400:00 | Yes
User Time | Extraction kicker trigger problem, there was a short circuit in PFN cable of the transformer. | PS |  | 2022-07-04 08:00 | 2022-07-04 12:10 | 4:10 | 250 | 14400:00 | No
User Time | Beam lost due to BPM interlock, libera#1 shows disconnection. | Diagnostic |  | 2022-07-05 06:22 | 2022-07-05 08:12 | 1:50 | 110 | 14400:00 | Yes
Beamline Preparation | MS beamline scientist asked to enter restricted access to SR with vacuum group to install rotating fillter | Front End |  | 2022-07-17 08:25 | 2022-07-17 17:50 | 9:25 | 565 | 14400:00 | No
User Time | Beam lost due to MPS communication fault | Control |  | 2022-07-20 13:04 | 2022-07-20 13:21 | 0:17 | 17 | 14400:00 | No
User Time | Water flow interlock at HESEB fixed mask at the front end | Cooling | Flow Switch | 2022-07-21 08:50 | 2022-07-21 09:17 | 0:27 | 27 | 14400:00 | No | 1 Min
Beamline Preparation | HESEB beamline scientist asked to enter restricted access to SR with Mechanical group for alignment | Front End |  | 2022-07-25 16:03 | 2022-07-25 19:27 | 3:24 | 204 | 14400:00 | No
User Time | Flow interlock at RF cavity#1 cooling rack | Cooling |  | 2022-07-26 08:25 | 2022-07-26 09:00 | 0:35 | 35 | 14400:00 | Yes
User Time | Heat absorbers interlock in C6 & C3, water flow interlock at HESEB front end, photon shutter & fixed mask. | Cooling |  | 2022-07-26 09:00 | 2022-07-26 11:27 | 2:27 | 147 | 14400:00 | No
User Time | Water flow interlock at HESEB front end, photon shutter & fixed mask. | Cooling |  | 2022-07-26 12:11 | 2022-07-26 12:46 | 0:35 | 35 | 14400:00 | Yes
Non-Interrupting | C7_SF1_Flow Interlock | Cooling |  | 2022-08-04 03:00 | 2022-08-04 08:56 | 5:56 | 356 | 14400:00 | Yes | Cause beam blowup
User Time | Beam lost due to MPS communication fault | Control |  | 2022-08-04 15:00 | 2022-08-04 15:30 | 0:30 | 30 | 14400:00 | No | 0 .. 1 no previous indication
User Time | leakage on cell #7 on the crotch absorber connection | Cooling |  | 2022-08-06 19:17 | 2022-08-07 08:30 | 13:13 | 793 | 14400:00 | No
Beamline Preparation | leakage on cell #7 on the crotch absorber connection | Cooling |  | 2022-08-06 19:19 | 2022-08-07 08:00 | 12:41 | 761 | 14400:00 | Yes
User Time | Microtron high voltage fuse was busted | PS |  | 2022-08-07 08:30 | 2022-08-07 09:30 | 1:00 | 60 | 14400:00 | No
User Time | Tuner of cavity#4 was not moving | RF |  | 2022-08-07 09:30 | 2022-08-07 10:35 | 1:05 | 65 | 14400:00 | No
Non-Interrupting | SR_VA_IMG_C14 | Vacuum |  | 2022-08-08 02:30 | 2022-08-08 02:30 | 0:00 | 0 | 14400:00 | No
Beamline Preparation | Beam lost due to water flowmeter interlock at amplifier#3 inlet | RF | SSA Flowmeter | 2022-09-19 10:26 | 2022-09-19 11:26 | 1:00 | 60 | 14400:00 | Yes
Non-Interrupting | Vacuum interlock at microtron | Vacuum | Microtron | 2022-09-20 08:00 | 2022-09-20 08:35 | 0:35 | 35 | 14400:00 | No
Beamline Preparation | Mechanical group entered SR to check the mechanical problem with undulator while opening/closing. | Front End |  | 2022-09-21 10:00 | 2022-09-21 16:00 | 6:00 | 360 | 14400:00 | No
Non-Interrupting | Vacuum interlock at microtron | Vacuum | Microtron | 2022-09-22 08:00 | 2022-09-22 08:35 | 0:35 | 35 | 14400:00 | No
Beamline Preparation | Mechanical group entered SR to check the mechanical problem with undulator while opening/closing. | Front End |  | 2022-09-22 10:00 | 2022-09-22 15:30 | 5:30 | 330 | 14400:00 | No
Non-Interrupting | AGPS Gun Spark interlock | PS |  | 2022-09-27 08:30 | 2022-09-27 08:40 | 0:10 | 10 | 14400:00 | No
User Time | Water flow interlock at MS front end | Cooling |  | 2022-09-28 08:48 | 2022-09-28 09:13 | 0:25 | 25 | 14400:00 | Yes | 20 min
Non-Interrupting | AGPS Gun Spark interlock | PS |  | 2022-10-03 08:15 | 2022-10-03 08:39 | 0:24 | 24 | 14400:00 | No
User Time | Beam lost due to wrong procedure while implementing the orbit bump | Human Error |  | 2022-10-03 08:58 | 2022-10-03 09:00 | 0:02 | 2 | 14400:00 | No
Non-Interrupting | AGPS Gun Spark interlock | PS |  | 2022-10-04 08:10 | 2022-10-04 08:22 | 0:12 | 12 | 14400:00 | No
Non-Interrupting | AGPS Gun Spark interlock | PS |  | 2022-10-05 08:20 | 2022-10-05 08:34 | 0:14 | 14 | 14400:00 | No
Non-Interrupting | AGPS Gun Spark interlock | PS |  | 2022-10-06 08:20 | 2022-10-06 08:34 | 0:14 | 14 | 14400:00 | No
User Time | Beam lost due to coupler#4 high temperature (Inlet/Outlet cooling pipes were closed) | Human Error |  | 2022-10-26 08:00 | 2022-10-27 08:00 | 24:00 | 1440 | 14400:00 | No
User Time | Electricity Dip | Power Grid |  | 2022-10-27 08:00 | 2022-10-28 08:00 | 24:00 | 1440 | 14400:00 | Yes
User Time | Beam lost due to cooling interlock at Cavity#4 flow switch | Cooling |  | 2022-10-30 03:50 | 2022-10-30 09:30 | 5:40 | 340 | 14400:00 | Yes
User Time | SR_VA_IMG_C14 | Vacuum |  | 2022-10-30 18:40 | 2022-10-30 18:50 | 0:10 | 10 | 14400:00 | No
User Time | AGPS Filament OverVoltage | PS | Aux. Gun | 2022-10-30 19:12 | 2022-10-30 19:30 | 0:18 | 18 | 14400:00 | No
User Time | SR_VA_IMG_C14 | Vacuum |  | 2022-10-31 06:44 | 2022-10-31 12:04 | 5:20 | 320 | 14400:00 | No
User Time | Beam lost due to Reflected power in Cavity 2&3 | RF |  | 2022-10-31 12:04 | 2022-10-31 13:00 | 0:56 | 56 | 14400:00 | No
User Time | Failure happened to the Cell06 QD1 power supply, we couldn’t cycle magnets in properly | PS |  | 2022-10-31 13:00 | 2022-10-31 14:00 | 1:00 | 60 | 14400:00 | No
User Time | SR_VA_IMG_C14 | Vacuum |  | 2022-11-01 06:40 | 2022-11-01 08:30 | 1:50 | 110 | 14400:00 | Yes
User Time | AGPS Filament OverVoltage | PS | Aux. Gun | 2022-11-01 08:30 | 2022-11-01 09:52 | 1:22 | 82 | 14400:00 | No
User Time | SR_VA_IMG_C14 | Vacuum |  | 2022-11-01 14:49 | 2022-11-01 15:14 | 0:25 | 25 | 14400:00 | Yes
User Time | SR_VA_IMG_C14 | Vacuum |  | 2022-11-02 19:46 | 2022-11-02 19:08 | -0.026389 | -38 | 14400:00 | Yes | Time entered maniually
User Time | XAFS Vacuum Interlock | Vacuum |  | 2022-11-03 00:18 | 2022-11-02 23:31 | -0.032639 | -47 | 14400:00 | Yes
User Time | AGPS Gun Spark interlock | PS |  | 2022-11-03 11:08 | 2022-11-03 11:18 | 0:10 | 10 | 14400:00 | No
User Time | Beam lost due to vacuum interlock at HESEB front end | Front End |  | 2022-11-03 11:50 | 2022-11-03 11:32 | -0.0125 | -18 | 14400:00 | Yes | HESEB new entry
User Time | Electricity Dip | Power Grid |  | 2022-11-03 22:45 | 2022-11-04 08:00 | 9:15 | 555 | 14400:00 | Yes
User Time | Beam lost when we turn the machine status to "User Time" instead of Injection. | Control |  | 2022-11-30 08:52 | 2022-11-30 09:47 | 0:55 | 55 | 14400:00 | No
User Time | Beam lost due to IR mirror thermocouples (MPS from BL side) | Front End |  | 2022-12-01 15:04 | 2022-12-01 15:08 | 0:04 | 4 | 14400:00 | Yes
User Time | Electricity Dip yesterday (3/12/2022) | Power Grid |  | 2022-12-04 08:15 | 2022-12-04 10:33 | 2:18 | 138 | 14400:00 | Yes
User Time | Beam lost due to known reason | Unknown |  | 2022-12-04 20:10 | 2022-12-04 20:41 | 0:31 | 31 | 14400:00 | Yes
User Time | Beam lost while doing bump for HESEB | Human Error |  | 2022-12-05 08:56 | 2022-12-05 09:00 | 0:04 | 4 | 14400:00 | No
User Time | Vacuum Interlock at microtron, which turned-off magnetron filament | Vacuum |  | 2022-12-06 08:00 | 2022-12-06 08:30 | 0:30 | 30 | 14400:00 | No
User Time | Microtron HV PS fuse busted | PS |  | 2022-12-06 08:30 | 2022-12-06 08:50 | 0:20 | 20 | 14400:00 | No
User Time | Beam lost due to thermo-couple interlock at RF cavities | RF |  | 2022-12-07 10:22 | 2022-12-07 10:40 | 0:18 | 18 | 14400:00 | Yes
User Time | Encountered saturation while injection, and we couldn't exceed 132mA, so we decided to kill the beam and re-inject again | Unknown |  | 2022-12-07 10:48 | 2022-12-07 10:52 | 0:04 | 4 | 14400:00 | No
User Time | Electricity Dip | Power Grid |  | 2022-12-14 11:45 | 2022-12-14 13:16 | 1:31 | 91 | 14400:00 | Yes
User Time | MS wiggler has High tilt alarm while opening | ID |  | 2022-12-19 08:05 | 2022-12-19 08:35 | 0:30 | 30 | 14400:00 | No
User Time | Electricity Dip | Power Grid |  | 2022-12-22 06:40 | 2022-12-22 10:08 | 3:28 | 208 | 14400:00 | Yes
User Time | Beam lost due to water flow interlock at SSA#3 | RF |  | 2022-12-25 13:09 | 2022-12-25 14:09 | 1:00 | 60 | 14400:00 | Yes
User Time | Beam lost due IR thermocouple#4 high temperature interlock | Front End |  | 2022-12-26 09:13 | 2022-12-26 09:32 | 0:19 | 19 | 14400:00 | Yes
User Time | Beam lost due to global interlock, it might be the "heseb_ff" script that killed the beam due to code bug | Human Error |  | 2022-12-26 09:52 | 2022-12-26 09:58 | 0:06 | 6 | 14400:00 | Yes
User Time | Electricity Dip | Power Grid |  | 2022-12-26 17:16 | 2022-12-26 18:40 | 1:24 | 84 | 14400:00 | Yes
User Time | Air flow interlock at RF cavities; due to solenoid valve wasn't opening (solenoid valve malfunction) | RF |  | 2022-12-27 00:09 | 2022-12-27 09:40 | 9:31 | 571 | 14400:00 | Yes
User Time | Cavity#2 tune error was high (about -70), it was hitting CW limit, and after moving some connection it showed CW & CCW limits together!! | RF |  | 2022-12-27 09:40 | 2022-12-27 10:08 | 0:28 | 28 | 14400:00 | No
User Time | Air flow interlock at RF cavities; due to solenoid valve wasn't opening (solenoid valve malfunction) | RF |  | 2022-12-27 10:59 | 2022-12-27 11:35 | 0:36 | 36 | 14400:00 | Yes
User Time | We struggle while injection, we inject 10mA and got saturation, we couldn't exceed it. We discovered that the BbB feedback system in X-Plane was running (It was not included in "Start_up_srrf" to turn-off the BbB in X) | Human Error |  | 2022-12-29 08:11 | 2022-12-29 09:01 | 0:50 | 50 | 14400:00 | No
User Time | Electricity Dip | Power Grid |  | 2022-12-29 18:27 | 2022-12-29 19:02 | 0:35 | 35 | 14400:00 | Yes
Injection | Water flow interlock at HESEB front end | Cooling |  | 2023-01-30 08:40 | 2023-01-30 08:50 | 0:10 | 10 | 14400:00 | Yes
Injection | One of the Emergency Push Buttons was hit accidentally, which leads to open SR tunnel (while entering restricted access to SR to solve previous problem) | Human Error |  | 2023-01-30 08:50 | 2023-01-30 09:24 | 0:34 | 34 | 14400:00 | No
Injection | Injection to SR was late due to AGPS interlock | Microtron |  | 2023-02-02 08:20 | 2023-02-02 08:37 | 0:17 | 17 | 14400:00 | No
Injection | Injection to SR was late due to problem while opening BEATS wiggler | ID |  | 2023-02-02 08:37 | 2023-02-02 09:04 | 0:27 | 27 | 14400:00 | No
Injection | There was a problem in closing MS wiggler; due to tilt error due to earthquake in Turkey !! | ID |  | 2023-02-06 08:50 | 2023-02-06 09:15 | 0:25 | 25 | 14400:00 | No
User Time | IMG1 vacuum spike interlock at MS front end (Fake Interlock) | Vacuum |  | 2023-02-08 00:05 | 2023-02-08 08:30 | 8:25 | 505 | 14400:00 | Yes
Injection | Injection was late due to aging of cathode, HV of AGPS increased to 2450V to overcome the problem temporary | Microtron |  | 2023-02-08 08:30 | 2023-02-08 09:12 | 0:42 | 42 | 14400:00 | No
Injection | Beam lost due to SSA#4 inlet water flow interlock, due to stuck flow turbine | RF |  | 2023-02-09 08:30 | 2023-02-09 09:53 | 1:23 | 83 | 14400:00 | Yes
Injection | AGPS trigger on Emission interlock | Microtron |  | 2023-02-18 08:38 | 2023-02-18 08:49 | 0:11 | 11 | 14400:00 | No
Injection | Microtron Dipole current was zero due to unknown reason | Microtron |  | 2023-02-20 08:00 | 2023-02-20 09:10 | 1:10 | 70 | 14400:00 | No
Injection | Microtron Dipole current was zero due to unknown reason | Microtron |  | 2023-02-21 08:00 | 2023-02-21 08:29 | 0:29 | 29 | 14400:00 | No
Injection | MS Wiggler didn't open due to different speed set-points between the motors | ID |  | 2023-03-08 08:15 | 2023-03-08 08:41 | 0:26 | 26 | 14400:00 | No
User Time | Water flow interlock at PNHL | Cooling |  | 2023-03-10 01:45 | 2023-03-10 08:00 | 6:15 | 375 | 14400:00 | Yes
User Time | Beam lost due to RF interlock (Water flow outlet AMP3) | RF |  | 2023-03-13 18:31 | 2023-03-13 20:13 | 1:42 | 102 | 14400:00 | Yes
User Time | Beam got lose due to a human mistake; instead of turning ON the MTBF for tune measurement, RF group operated it as feedback by mistake which made a beam blowup. | Human Error |  | 2023-03-14 15:08 | 2023-03-14 15:20 | 0:12 | 12 | 14400:00 | Yes
Injection | SR Libra#7 IOC hanged, so the correction was not running in a correct way | Diagnostic |  | 2023-03-23 09:30 | 2023-03-23 10:26 | 0:56 | 56 | 14400:00 | Yes
Injection | Trigger sequencer was not triggering-on the AGPS | Control |  | 2023-04-04 08:10 | 2023-04-04 08:45 | 0:35 | 35 | 14400:00 | No
User Time | Beam got lost after enabling BPM interlock (Due to bpm_default command it returned the limit values of HESEB BPM to 1mm instead of 1.5mm)) | Human Error |  | 2023-04-04 09:12 | 2023-04-04 09:20 | 0:08 | 8 | 14400:00 | Yes
User Time | Water flow interlock at MS beamline rotation filter | Cooling |  | 2023-04-04 10:02 | 2023-04-04 10:33 | 0:31 | 31 | 14400:00 | No
User Time | Water flow interlock at MS beamline rotation filter | Cooling |  | 2023-04-05 10:39 | 2023-04-05 11:02 | 0:23 | 23 | 14400:00 | No
User Time | Beam got lost due to IR THC#4 interlock; due to C11-QD1-PS current was decreasing | PS | QD | 2023-04-06 23:30 | 2023-04-07 10:30 | 11:00 | 660 | 14400:00 | Yes
Injection | Electricity Cut | Power Grid |  | 2023-04-07 10:30 | 2023-04-07 14:56 | 4:26 | 266 | 14400:00 | Yes
User Time | Heat absorber interlock at (SRC08-VA-HAB2-FLS2) | Cooling |  | 2023-04-10 04:23 | 2023-04-10 08:00 | 3:37 | 217 | 14400:00 | Yes
Injection | Microtron Dipole current was zero due to unknown reason | Microtron |  | 2023-04-12 08:20 | 2023-04-12 08:45 | 0:25 | 25 | 14400:00 | No
Injection | The beam lost while ramping (2.3 Gev) due to vacuum interlock at SR (SRC14-VA-IMG1); most probably we forgot to turn-on the orbit correction or to include RF Frequency option in correction. | Human Error |  | 2023-04-16 08:31 | 2023-04-16 08:52 | 0:21 | 21 | 14400:00 | Yes
Injection | Beam lost due to RF amplifier 3 water flow outlet interlock | RF |  | 2023-05-09 08:07 | 2023-05-09 08:17 | 0:10 | 10 | 14400:00 | Yes
User Time | Beam lost due to RF amplifier 3 REV Power interlock, Spare LLE problem | RF |  | 2023-05-09 08:45 | 2023-05-09 09:21 | 0:36 | 36 | 14400:00 | Yes
User Time | Water flow interlock at ID11 shutter | Cooling |  | 2023-05-09 15:13 | 2023-05-09 15:40 | 0:27 | 27 | 14400:00 | Yes
User Time | Beam lost due to RF amplifier 3 REV Power interlock, Spare LLE problem | RF |  | 2023-05-10 18:22 | 2023-05-10 18:33 | 0:11 | 11 | 14400:00 | Yes
User Time | Beam lost due to RF amplifier 3 REV Power interlock, Spare LLE problem | RF |  | 2023-05-10 21:20 | 2023-05-10 21:35 | 0:15 | 15 | 14400:00 | Yes
User Time | Beam lost due to RF amplifier 3 REV Power interlock, Spare LLE problem | RF |  | 2023-05-11 18:00 | 2023-05-11 19:20 | 1:20 | 80 | 14400:00 | Yes
User Time | Beam got lost due to high temperature at IR THC#4; due to larger vertical beam size 90um, instead of 87um | Human Error |  | 2023-05-14 09:24 | 2023-05-14 09:38 | 0:14 | 14 | 14400:00 | Yes
Injection | Beam got lost when running FF for BEATS after closing it to 30 (Human Error), due to distortion of orbit which exceeded the BPM interlock limits | Human Error |  | 2023-05-14 10:03 | 2023-05-14 10:08 | 0:05 | 5 | 14400:00 | Yes
User Time | Beam lost due to RF amplifier 3 water flow outlet interlock | RF |  | 2023-05-18 15:03 | 2023-05-18 15:35 | 0:32 | 32 | 14400:00 | Yes
User Time | Beam got lost due to IR THC#4 interlock, even there was no temperature rising, the root cause still unknown | Unknown |  | 2023-05-18 19:59 | 2023-05-18 20:11 | 0:12 | 12 | 14400:00 | Yes
Injection | Beam lost due to human error; the BPM interlock got enabled before run "beats_ff" | Human Error |  | 2023-05-21 08:38 | 2023-05-21 08:44 | 0:06 | 6 | 14400:00 | Yes
User Time | Beam got lost due to IR THC#4 interlock, even there was no temperature rising, the root cause still unknown | Unknown |  | 2023-05-23 15:57 | 2023-05-23 16:10 | 0:13 | 13 | 14400:00 | Yes
Injection | Booster RF Interlock | RF | Booster | 2023-06-06 08:12 | 2023-06-06 08:14 | 0:02 | 2 | 14400:00 | No
Injection | Booster RF Interlock | RF | Booster | 2023-06-07 08:08 | 2023-06-07 08:30 | 0:22 | 22 | 14400:00 | No
Injection | SR Cell-14 Vacuum interlock, the master oscillator didn't response to orbit correction (might be a human error) | Human Error |  | 2023-06-07 08:41 | 2023-06-07 09:14 | 0:33 | 33 | 14400:00 | Yes
Injection | Booster RF Interlock | RF | Booster | 2023-06-08 08:15 | 2023-06-08 08:16 | 0:01 | 1 | 14400:00 | No
User Time | Beam got lost due to IR THC#4 interlock | Unknown |  | 2023-06-08 17:22 | 2023-06-08 17:52 | 0:30 | 30 | 14400:00 | Yes
User Time | Electricity Cut | Power Grid |  | 2023-06-09 06:31 | 2023-06-09 12:16 | 5:45 | 345 | 14400:00 | Yes
Injection | We encountered a problem while closing the MS wiggler due to power cut happened on Friday, solved by control group | Power Grid | Wiggler | 2023-06-11 08:40 | 2023-06-11 09:00 | 0:20 | 20 | 14400:00 | Yes
Injection | Booster RF Interlock | RF | Booster | 2023-06-12 08:14 | 2023-06-12 08:15 | 0:01 | 1 | 14400:00 | No
Injection | Booster RF Interlock | RF | Booster | 2023-06-13 08:31 | 2023-06-13 08:33 | 0:02 | 2 | 14400:00 | No
Injection | Beam lost due to RF interlock; still under investigation. | RF |  | 2023-06-13 08:37 | 2023-06-13 08:46 | 0:09 | 9 | 14400:00 | Yes
Injection | Booster RF Interlock | RF | Booster | 2023-06-14 08:19 | 2023-06-14 08:36 | 0:17 | 17 | 14400:00 | No
Injection | Beam lost due to Quadrupole current deviation interlock; it seems related to a communication error, not to a real deviation (gateWay). | Control |  | 2023-06-15 08:11 | 2023-06-15 08:17 | 0:06 | 6 | 14400:00 | Yes
Injection | Booster RF Interlock | RF | Booster | 2023-06-15 08:22 | 2023-06-15 08:29 | 0:07 | 7 | 14400:00 | No
Injection | Beam lost while ramping due to RF ref.power | RF |  | 2023-06-18 08:31 | 2023-06-18 08:35 | 0:04 | 4 | 14400:00 | Yes
Injection | Beam lost while ramping due to RF ref.power | RF |  | 2023-06-19 08:37 | 2023-06-19 08:42 | 0:05 | 5 | 14400:00 | Yes
User Time | Electricity dip | Power Grid |  | 2023-06-19 14:24 | 2023-06-19 14:37 | 0:13 | 13 | 14400:00 | Yes
Injection | Beam lost while ramping (RF vacuum interlock) | RF |  | 2023-06-21 08:40 | 2023-06-21 08:44 | 0:04 | 4 | 14400:00 | Yes
Injection | Booster RF Interlock | RF | Booster | 2023-06-21 08:20 | 2023-06-21 08:22 | 0:02 | 2 | 14400:00 | No
User Time | Electricity dip | Power Grid |  | 2023-06-22 05:19 | 2023-06-22 08:05 | 2:46 | 166 | 14400:00 | Yes
Injection | Beam lost due to beam blowing-up while ramping (the correction was working fine, but we noticed that it was slower than it used to be | Unknown |  | 2023-07-17 08:33 | 2023-07-17 08:41 | 0:08 | 8 | 14400:00 | Yes
Injection | Quadrant#3 PS gateway was not responding , so the control group replaced it with another one. | Control |  | 2023-07-20 08:15 | 2023-07-20 10:18 | 2:03 | 123 | 14400:00 | Yes
Injection | ID 10 was not open | ID | ID10 | 2023-07-21 08:00 | 2023-07-21 09:37 | 1:37 | 97 | 14400:00 | No
User Time | Beam lost due to SCR02 QD1 & QD2 went to zero current, so after troubleshooting we found out that there was a phase failure (one phase missing). | Electrical | QD | 2023-07-21 22:35 | 2023-07-22 08:00 | 9:25 | 565 | 14400:00 | Yes
Machine Time | While doing cycling, SCR02 QD1 & QD2 was giving zero Ampere, so after trouble shooting we found out that there was a phase failure (one phase missing), then power supply group came and changed the burned cables. | PS | QD | 2023-07-22 08:00 | 2023-07-22 19:20 | 11:20 | 680 | 14400:00 | Yes
Non-Interrupting | Microtron magnet interlock due to power dip | Power Grid |  | 2023-07-23 17:08 | 2023-07-23 17:58 | 0:50 | 50 | 14400:00 | Yes
Injection | The beam late due to ID10 was not open | ID |  | 2023-07-24 08:15 | 2023-07-24 09:38 | 1:23 | 83 | 14400:00 | No
Injection | Beam lost while ramping, and there was saturation while injection. | Unknown |  | 2023-07-30 08:31 | 2023-07-30 08:56 | 0:25 | 25 | 14400:00 | Yes
User Time | Beam lost due to cooling chillers turned off (both of them) | Cooling |  | 2023-07-30 21:22 | 2023-07-31 00:09 | 2:47 | 167 | 14400:00 | Yes
User Time | We lost the beam because one of the chillers tripped, chiller flow switch failure. | Cooling |  | 2023-07-31 01:08 | 2023-07-31 08:20 | 7:12 | 432 | 14400:00 | Yes
Non-Interrupting | AGPS HV OverCurrent interlock, the cause still unknown; but sofian tighten some HV cable connection (Cause injection late, but not affecting on beam availability) | PS |  | 2023-08-01 08:10 | 2023-08-01 08:35 | 0:25 | 25 | 14400:00 | No
Non-Interrupting | AGPS HV OverCurrent interlock, the cause still unknown (Cause injection late, but not affecting on beam availability) | PS |  | 2023-08-01 10:20 | 2023-08-01 10:40 | 0:20 | 20 | 14400:00 | No
User Time | Beam lost due to water flow inlet interlock at amplifier#2. | RF | SSA Flowmeter | 2023-08-01 10:27 | 2023-08-01 11:08 | 0:41 | 41 | 14400:00 | Yes
Injection | Beam lost while ramping due to unknown cause (might related to RF ramping curve) | RF |  | 2023-08-02 08:39 | 2023-08-02 08:49 | 0:10 | 10 | 14400:00 | Yes
Machine Time | SR BM problem, BWM fiber receiver found faulty which caused IGBT damaged and output filet damaged as well. | PS | SR_BM | 2023-09-03 08:00 | 2023-09-07 18:00 | 106:00 | 6360 | 14400:00 | Yes
Machine Time | Beam lost due to Cooling interlock | Cooling | Flow Switch | 2023-09-07 21:42 | 2023-09-08 09:41 | 11:59 | 719 | 14400:00 | Yes
User Time | Beam lost due to SR-PS-SD1 earth leakage interlock (The water conductivity was high; due to bad water quality, which increased the Iearth) | Cooling |  | 2023-09-10 23:03 | 2023-09-10 23:23 | 0:20 | 20 | 14400:00 | Yes
Non-Interrupting | HV over current fault (happened twice); PS group changed control cable between PS and the controller. | PS |  | 2023-09-11 08:20 | 2023-09-11 08:42 | 0:22 | 22 | 14400:00 | Yes
User Time | Beam lost due to earth leakage in SD1 sextupole. It's due to cooling interlock (Water conductivity was high). | Cooling |  | 2023-09-11 16:40 | 2023-09-11 16:48 | 0:08 | 8 | 14400:00 | Yes
User Time | Beam lost due to quadrupoles current deviation interlock due to SRC15-PS-QF1 went to zero instantaneously, then came back to normal value (gateWay). | Control | EPICS gate way | 2023-09-12 01:36 | 2023-09-12 08:24 | 6:48 | 408 | 14400:00 | Yes
User Time | Beam lost due to sextupoles current deviation due to SD PS earth leakage interlock (The water conductivity was high; due to bad water quality, which increased the Iearth) | Cooling |  | 2023-09-12 20:27 | 2023-09-12 20:46 | 0:19 | 19 | 14400:00 | Yes
Non-Interrupting | Vacuum interlock in microtron | Vacuum | Microtron | 2023-09-12 23:30 | 2023-09-13 08:14 | 8:44 | 524 | 14400:00 | No
User Time | Beam lost due to SRC06-QF1 cooling flow interlock. | Cooling | Flow Switch | 2023-09-13 15:43 | 2023-09-13 15:52 | 0:09 | 9 | 14400:00 | Yes
Non-Interrupting | There was a vacuum interlock at microtron (It seems a fake interlock). | Vacuum | Microtron | 2023-09-13 23:40 | 2023-09-14 08:36 | 8:56 | 536 | 14400:00 | No
Non-Interrupting | There was a vacuum interlock at microtron its seems the fault is repeatly every night at the same time between ( 23 :20 to 23:45 ) and its a fake fault . | Vacuum | Microtron | 2023-09-14 23:44 | 2023-09-15 08:13 | 8:29 | 509 | 14400:00 | No
User Time | Beam lost due to RF cavity#4 out air interlock (It seems a fake interlock);Control Team increase the time of cavity 4 air out temp interlock to 15 sec | RF | Cavity4 | 2023-09-16 20:55 | 2023-09-17 07:51 | 10:56 | 656 | 14400:00 | Yes
Injection | Beam loast while ramping due to RF cooling rack thermocouple swaped without change the set points | RF | Cavity4 | 2023-09-18 09:01 | 2023-09-18 09:21 | 0:20 | 20 | 14400:00 | Yes
Injection | Beam lost While ramping due to RF ramping not work | Unknown | Unknown | 2023-09-18 09:21 | 2023-09-18 09:27 | 0:06 | 6 | 14400:00 | Yes
User Time | Beam lost due to problem in the dipole and vertical corrector power supplies gateway. | Control | EPICS gate way | 2023-09-20 17:43 | 2023-09-20 17:49 | 0:06 | 6 | 14400:00 | Yes
User Time | Beam lost due to SR-PS-SD1 SO interlock | PS |  | 2023-09-21 10:28 | 2023-09-21 10:33 | 0:05 | 5 | 14400:00 | Yes
User Time | Beam got lost due to C3-RF-HAB3 cooling interlock (This is perdicted to happen after the water replacement due to air bubbles | Cooling | RF Secondary Pumps | 2023-09-23 03:21 | 2023-09-23 07:57 | 4:36 | 276 | 14400:00 | Yes
User Time | SRC6-CO-QF1 flow interlock; due to defected magnetic switch, control & cooling groups entered SR restricted access and replaced it with new one. | Cooling | Flow Switch | 2023-09-26 13:20 | 2023-09-26 14:23 | 1:03 | 63 | 14400:00 | Yes
Non-Interrupting | Microtron vacuum interlock (Seems to be fake), and wave-guide pressure interlock. | Vacuum | Microtron | 2023-09-26 23:26 | 2023-09-27 08:39 | 9:13 | 553 | 14400:00 | No
User Time | Beam lost due to quadruple current deviation interlock; but unfortunately there was no indication at all on the archiver that shows any deviation in any of the quadruples !! (gateWay) | Control | EPICS gate way | 2023-10-02 22:22 | 2023-10-02 23:02 | 0:40 | 40 | 14400:00 | Yes
Non-Interrupting | Vacuum interlock in the Micotron : Its seems fake fault because the pressure in micotron wasn't changed at all | Vacuum | Microtron | 2023-10-02 23:48 | 2023-10-03 08:03 | 8:15 | 495 | 14400:00 | No
User Time | Beam lost due to fault on SRC15-PS-QD1 its seems fake fault too because the current decreasing to Zero and go back to normal only in 1 second (gateWay). | Control | EPICS gate way | 2023-10-03 06:15 | 2023-10-03 08:15 | 2:00 | 120 | 14400:00 | Yes
User Time | Beam lost due to Quadruples interlock current deviation interlock (SRC09-QF1) there was no indication on PS GUI! (gateWay). | Control | EPICS gate way | 2023-10-03 19:59 | 2023-10-03 20:52 | 0:53 | 53 | 14400:00 | Yes
User Time | Beam Lost due to an issue in rf amplifier 4 cooling | RF | Cavity4 | 2023-10-05 04:48 | 2023-10-05 09:20 | 4:32 | 272 | 14400:00 | Yes
User Time | Beam lost due to quadrupols current deviation interlock due to SRC15-PS-QD1 current went to zero instantaneously, then came back to normal value (gateWay). | Control | EPICS gate way | 2023-10-08 15:15 | 2023-10-08 15:32 | 0:17 | 17 | 14400:00 | Yes
Injection | Injection was late because of the control Team Replaced the gateway 3-Q4 | Control | EPICS gate way | 2023-10-09 08:15 | 2023-10-09 09:06 | 0:51 | 51 | 14400:00 | No
User Time | Quadrupoles current deviation, there was no indication on PS GUI! (gateWay). | Control | EPICS gate way | 2023-10-09 09:32 | 2023-10-09 09:42 | 0:10 | 10 | 14400:00 | Yes
User Time | Beam lost Due to Electrical failure on SR main circuit Breaker | Electrical |  | 2023-10-10 06:47 | 2023-10-10 13:56 | 7:09 | 429 | 14400:00 | No
User Time | Beam lost due to RF failure delta water flow on amplifier 4 , the RF Team changed the flow meter to slove this problem | RF | SSA Flowmeter | 2023-10-11 09:05 | 2023-10-11 09:23 | 0:18 | 18 | 14400:00 | Yes
User Time | Beam lost due to RF failure water flow outlet on amplifier 4 | RF | SSA Flowmeter | 2023-10-11 10:14 | 2023-10-11 10:32 | 0:18 | 18 | 14400:00 | Yes
User Time | Beam lost due to motion Team failure | ID | ID10 | 2023-10-11 10:54 | 2023-10-11 11:08 | 0:14 | 14 | 14400:00 | Yes
Injection | Beam Lost while Ramping its seems a correction fault | RF | Cavity3 | 2023-10-11 09:04 | 2023-10-11 08:43 | -0.014583 | -21 | 14400:00 | Yes
Injection | Beam Lost while Ramping its seems a correction fault | RF | Cavity3 | 2023-10-11 10:15 | 2023-10-11 08:43 | -0.063889 | -92 | 14400:00 | Yes
Injection | Beam Lost while Ramping its seems a correction fault | RF | Cavity3 | 2023-10-11 10:55 | 2023-10-11 08:43 | -0.091667 | -132 | 14400:00 | Yes
User Time | Cooling issue in MS beam line front End | Cooling | Flow Switch | 2023-10-18 09:36 | 2023-10-18 10:00 | 0:24 | 24 | 14400:00 | Yes
Injection | Beam lost while Ramping | RF | Cavity3 | 2023-10-19 08:33 | 2023-10-19 08:40 | 0:07 | 7 | 14400:00 | Yes
Injection | Libra 1 was down due to the power dip happened twice | Power Grid |  | 2023-10-22 08:15 | 2023-10-22 09:03 | 0:48 | 48 | 14400:00 | Yes
Injection | we killed the Beam because of a problem in the correction | Human Error |  | 2023-10-22 09:21 | 2023-10-22 09:25 | 0:04 | 4 | 14400:00 | Yes
User Time | Beam lost due to spark in Cavity#4 | RF |  | 2023-10-22 22:41 | 2023-10-23 08:43 | 10:02 | 602 | 14400:00 | Yes
Injection | Beam lost During Ramping RF cavity #3 amplifier, RF group changed cooling rack temperature set-point from 62C to 61.5C. | RF | Cavity3 | 2023-10-25 08:35 | 2023-10-25 08:40 | 0:05 | 5 | 14400:00 | Yes
Injection | Beam lost while ramping at 2.0GeV, because when we run ramping script, it opened a matlab session and for the first time it requests a password, and we didn't notice it, so the magnets start to ramp but RF wasn't, so the beam get lost. | Human Error |  | 2023-11-12 08:34 | 2023-11-12 08:44 | 0:10 | 10 | 14400:00 | Yes
Injection | Beam lost while ramping due to RF interlock | RF |  | 2023-11-16 08:33 | 2023-11-16 08:40 | 0:07 | 7 | 14400:00 | Yes
Injection | Beam lost suddenly while ramping within one minutes from 250mA to 40mA; with no clear reason. | Unknown |  | 2023-11-19 08:50 | 2023-11-19 08:54 | 0:04 | 4 | 14400:00 | Yes
Injection | Trigger On Emission Fault due to Microtron dipole current was zero (Dipole was off) due to unknown reason. | Microtron |  | 2023-11-20 09:00 | 2023-11-20 09:10 | 0:10 | 10 | 14400:00 | No
Injection | Booster extraction kicker was not working, the power supplies group found the base of Thyrotron (in Booster Extraction kicker )is damaged so PS group changed it with a new base and thyrotron. | PS |  | 2023-11-20 09:10 | 2023-11-20 15:13 | 6:03 | 363 | 14400:00 | No
Injection | Wave-guide pressure interlock | RF |  | 2023-12-03 08:15 | 2023-12-03 08:25 | 0:10 | 10 | 14400:00 | No
User Time | Beam lost due to Cavity2 Tuner Motion Box Stuck | RF |  | 2023-12-03 22:22 | 2023-12-04 08:17 | 9:55 | 595 | 14400:00 | Yes
User Time | Electricity dip | Power Grid |  | 2023-12-05 22:23 | 2023-12-06 08:00 | 9:37 | 577 | 14400:00 | Yes
Injection | Beam lost due to electricity fluctation, Which cause SR-BM to get damaged (fixed by changing some electrocnics boards and resistors) | PS |  | 2023-12-05 22:23 | 2023-12-07 16:15 | 41:52 | 2512 | 14400:00 | Yes
User Time | Beam lost due to SRC15-BPM2 is interlocked (LBR12-BPM1); the BPM module is replaced with spare one | Diagnostic |  | 2023-12-10 15:27 | 2023-12-10 17:01 | 1:34 | 94 | 14400:00 | Yes
User Time | Beam lost due to dipole interlock due to Chiller3 interlock (Freeze protection) | Cooling |  | 2023-12-16 15:19 | 2023-12-16 15:45 | 0:26 | 26 | 14400:00 | Yes
User Time | Beam lost due Cooling Rack #2 had lost 3 3Phase Contactors (burned out); Pump, Heater1 and Heater2. RF group changed it. | RF |  | 2023-12-22 07:06 | 2023-12-23 08:00 | 24:54 | 1494 | 14400:00 | Yes
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
[6 empty] | 0:00 | 0 | 14400:00
Note on B14: Author: Sudden drop in the quadropoles current in SR.
Note on B20: Author: BEAM LOST DUE TO STEP TUNE BEFORE CORRECT OPTICS
Note on B23: Author: beam lost at Energy 800 Mev
Note on B26: Author: beam lost at Energy 800 Mev
Note on B33: Author: SR bending magnet couldn't be cycled; due to communication error.
Note on B37: Author: Beam lost due to BPM interlock, the root cause is unknown
Note on B44: Author: and HAB3
Note on B49: Author: Electricity Fluctuation Three of PLC's UPS have been failed SR Dipole & RF power supply main circuit breakers were tripped Microtron magnet tripped Cooling fault at SR C06 QD2 & HAB3 due to flow switch stuck up
Note on B54: Author: and HAB3
Note on B58: Author: Short circuit happened on Triplet #1 QD which cut off the power on the Trim coils, TL1 correctors and the microtron vacuum gauge
Note on H76: Author: Exelcuding the non-user time and holidays
Note on B77: Author: RF tripped while ramping, due to EPICS gate way didn't give permission to Matlab to write on the LLEs, you can see below that the RF cavities voltage didn't change while ramping.
Note on B78: Author: RF tripped while ramping, due to EPICS gate way didn't give permission to Matlab to write on the LLEs, you can see below that the RF cavities voltage didn't change while ramping.
Note on B79: Author: Beam lost; due to BPM interlock, after ramping, whenever We enable the BPM interlock, we lose the beam.
Note on B85: Author: but it doesn't show any interlock related to power supply, it just shows RF interlock
Note on B86: Author: but it doesn't show any interlock related to power supply, it just shows RF interlock
Note on B87: Author: but it doesn't show any interlock related to power supply, it just shows RF interlock
Note on B88: Author: but it doesn't show any interlock related to power supply, it just shows RF interlock
Note on B90: Author: The reason of the fault is still unknown but the power Supply group is working on it to find the root cause of the trip. The power supply group did change the crowbar of dipole power supply.
Note on B92: Author: it shows on the screen that it's heating-up, but when we check the cabinet, we see that the Aux. Gun voltage rising up till 2.3 Kv, but the filament current still zero. The problem was due to the filament potentiometer current was set to zero by mistake.
Note on B95: Author: PCO interlock
Note on B98: Author: C13_QF1&QF2 C13_QD1&QD2 C14_QF1 C15_QF1 C16_QF2
Note on B100: Author: Beam was't lost, it didn't affect the beam. It was in user time, but I put it in decay because the beam wan't affected.
Note on B101: Author: there was no actual interlock either in storage ring IMG1 nor in front end IMG1
Note on B102: Author: Control group changed the Power supply of the booster RF control signals.
Note on B104: Author: orbit correction was not working at top energy.
Note on B106: Author: IEarth ILK in the Cell 04 QD2 Power Supply.
Note on B107: Author: We noticed a huge phase error oscillation on cavity4 at the time of beam lost, we reset the RF system, and tried to start the RF, but We failed; due to the big phase error in cavity3&4. We contacted Darweesh, and with his help, we reduced the phase error in both cavities, and started RF again.
Note on B110: Author: happened 6 times
Note on B114: Author: While ramping
Note on B115: Author: While ramping
Note on B119: Author: A huge saturation while injection, we killed the beam and reinject again.
Note on B120: Author: Changing the aux. gun didn't take too much time, but due to vacuum leake we took about three days.
Note on B121: Author: Changing the aux. gun didn't take too much time, but due to vacuum leake we took about three days.
Note on B125: Author: loose connection of a cable in the PLC communication module
Note on B126: Author: loose connection of a cable in the PLC communication module
Note on B127: Author: loose connection of a cable in the PLC communication module
Note on B134: Author: Beam lost due to RF trip while ramping
Note on B135: Author: Beam lost due to RF trip while ramping
Note on B138: Author: Beam lost due to RF trip while ramping
Note on B139: Author: Beam lost due to RF trip while ramping
Note on B140: Author: Beam lost due to RF trip while ramping
Note on B142: Author: Beam lost due to RF trip while ramping
Note on B150: Author: We consider it in decay mode, because it didn't disturbe the beam.
Note on B151: Author: We consider it in decay mode, because it didn't disturbe the beam.
Note on B153: Author: control group and Yazeed entered the SR tunnel to fix and calibrate the Wiggler.
Note on B157: Author: The lost due to the human error in material science beamline by accidentally turning off some vacuum gauges.
Note on B158: Author: control group and Yazeed entered the SR tunnel to fix and calibrate the Wiggler.
Note on B160: Author: Beam lost while ramping due to RF voltage stopped increasing with SR-BM current, after overcoming the phase errors issue in RF cavities. Ramping script need to be reviewed according to RF group feedback.
Note on B166: Author: Gihan turned on the circuit breaker of the experimental hutch at that morning, and it tripped.
Note on B174: Author: While Ramping
Note on B177: Author: The beam lost due to cooling and its human fault since i didn't notice cooling alarm.
Note on B178: Author: While Ramping
Note on B179: Author: Beam lost due to fake alarm of cooling system high temperature on QD2 in cell#9.
Note on B180: Author: RF Cavity#2 Limit switch CCW in LLE1 were active even though the limit switch is not hit.
Note on B182: Author: Libera4-BPM disconnection during injection, which stopped orbit correction and big shift in master oscillator frequency, and beam lost!
Note on B185: Unknown global interlock
Note on B187: Overload on the usage of memory of IO. The register GUI of AGPS was open twice and there was a huge request on IOCs, which consume 100% of IOC memory usage. One of them has been closed, and the problem has been solved.
Note on B188: Beam lost while ramping due to unknown global interlock.
Note on B189: Unknown global interlock
Note on B190: Beam lost while ramping due to something wrong with tuner motion box of RF cavity (Big difference in klystron phase)
Note on B191: Beam lost due to RF at full Energy (Big difference in klystron phase)
Note on B194: Beam lost while ramping due to unknown global interlock.
Note on B196: Unknown global interlock
Note on B197: Unknown global interlock
Note on B201: Beam lost due to RF at full Energy (Big difference in klystron phase)
Note on B203: Beam lost due to RF while ramping (Big difference in klystron phase) Tuner problem
Note on B204: Darweesh saw sharp turbulence in water flow to amplifiers
Note on B215: While Ramping
Note on B216: While Ramping
Note on B220: there was a huge instability in "Y" plane; close to vertical tune
Note on B234: it gives interlock even the interlock cable disconnected, After this interlock we decided to change the LLE.
Note on B244: starts to appear lower than 150mA
Note on B246: One of gauge controllers was disconnected
Note on B249: even though there was no alarm on the correction matlab session !
Note on B251: We faced a high tuner error for cavity#2, which took time until return it manually to the acceptable range.
Note on B252: Mahmoud AbdLatif and Yazeed asked for restricted access to the SR
Note on B265: We couldn’t injection into SR
Note on B268: communication in MPS PLC which caused all SR-SHUTTERS closed and kill the RF
Note on B275: PLC communication Error in the IMGs which cause the MPS interlock in the PLC ,communication module was replaced by the Control Group.
Note on B276: Someone hit accidentally
Note on B277: Someone hit accidentally
Note on B279: So RF group entered to SR in restricted mode and fix it
Note on B296: Fake Vacuum interlock, from IMG channel
Note on B298: Fake Vacuum interlock, from IMG channel
Note on B301: Fake Vacuum interlock, from IMG channel
Note on B303: Fake Vacuum interlock, from IMG channel
Note on B304: Fake Vacuum interlock, from IMG channel
Note on B309: It's something related to some new changed and modifications on control system.
Note on B323: See the sharp edge of QF1 cell#11 while ramping
Note on B339: AGPS trigger on Emission interlock due to microtron dipole was off (the current was zero without any indication)
Note on B340: AGPS trigger on Emission interlock due to microtron dipole was off (the current was zero without any indication)
Note on B341: (one of the motors speed was 1mm/s, the other was 0.5mm/s);so the motor open with tilt correction speed which is 0.1mm/s, this was due to the power cut on last Sunday.
Note on B345: Beam got lost while ramping due to vacuum interlock in cell#9
Note on B348: so MS beamline scientist request to kill the beam to enter restricted access to SR
Note on B349: so MS beamline scientist request to kill the beam to enter restricted access to SR
Note on B353: AGPS trigger on Emission interlock due to microtron dipole was off (the current was zero without any indication)
Note on B361: it might be related to not run FF for HESEB while it's closing
Note on B371: BbB feedback ??
Note on B441: RF Comment: no specific misbehavior were seen in the LLE pv's and most probably the trip happened due to driven instability in the machine.

## Master
Trips | Machine Modes | RF | PS | Vacuum | Cooling | Control | Diagnostic | Unknown | Human Error | Power Grid | Front End | Network | ID | Magnet | Microtron | Electrical
RF | Injection | LLE1 | Microtron Magnat | Ion Pump Controller | Chillers Pumps | VME Crate | Beam Position Monitor | Unknown |  | Power Grid Dip | Pneumatic | Server | ID09
PS | User Time | LLE2 | Modulator | Ion Pump Splitter | Chillers | Moxa Terminal Server | Beam Current Monitor | During Ramp |  |  |  |  | ID10
Vaccum | Machine Time | Signal Gen. | Magnetron | Ion Pump | Heat Exchangers | Programmable Logic Controller | Faraday Cup | [5 empty] | ID11
Cooling | Non-Interrupting | SSA1 | Aux. Gun | Gauge Controller | Primary | Power Supplies | Foil Monitor
Control | Free | SSA2 | Trim Coil | Pirani Gauge | SR Secondary Pumps | Communication error | Libera 1
Diagnostic | Beamline Preparation | SSA3 | Triplet | Inverted Magnetron Gauge | RF Secondary Pumps | EPICS gate way | Libera 2
Human Error |  | SSA4 | SR_BM | Gate Valve | BO Secondary Pumps |  | Libera 3
Unknown |  | Cavity1 | BO_BM | Residual Gas Analyzer | BL Secondary Pumps |  | Fast Current Transformer
Electricity |  | Cavity2 | QF | Titanium Sublimation Pump Controller | Flow Switch |  | New Parametric Current Transformer
|  | Cavity3 | QD | Titanium Sublimation Pump | Heat Absorber |  | Direct Current Current Transformer
|  | Cavity4 | SF | Fast Valve Controller |  |  | Camera
|  | Booster | SD | Fast Valve |  |  | Shaker
|  | Klystron Phase Correction | HC | Microtron |  |  | Multiplexer
|  | SSA Flowmeter | VC |  |  |  | Signal Switcher
|  |  | SQ |  |  |  | Fast Interlock
BO_INJ_SP
BO_INJ_KC
BO_EX_SP
BO_EX_KC
SR_INJ_SP
SR_INJ_KC
PCO
QF & QD